## Slide 1: T24_vg1.8_disk 11WNSDVG1.8 CLIC_G un-damped @ 11.424 GHz measurements versus simulations
Grudiev
CERN
1.07.2009

## Slide 2: Acknowledgements
CERN:
M. Gerbaux
R. Zennaro
A. Olyunin
W. Wuensch
SLAC:
Z. Li

## Slide 3: First cell
Hs/Ea
Es/Ea
Sc/Ea2
a [mm] | 3.307
d [mm] | 1.753
e | 1.16
f [GHz] | 11.433
Q(Cu) | 6814
vg/c [%] | 1.83
r’/Q [LinacΩ/m] | 15198
Es/Ea | 1.95
Hs/Ea [mA/V] | 2.6
Sc/Ea2 [mA/V] | 0.37

## Slide 4: Middle cell
Hs/Ea
Es/Ea
Sc/Ea2
a [mm] | 2.887
d [mm] | 1.402
e | 1.15
f [GHz] | 11.432
Q(Cu) | 6980
vg/c [%] | 1.33
r’/Q [LinacΩ/m] | 16960
Es/Ea | 1.95
Hs/Ea [mA/V] | 2.45
Sc/Ea2 [mA/V] | 0.34

## Slide 5: Last cell
Hs/Ea
Es/Ea
Sc/Ea2
a [mm] | 2.467
d [mm] | 1.05
e | 1.13
f [GHz] | 11.426
Q(Cu) | 7157
vg/c [%] | 0.92
r’/Q [LinacΩ/m] | 18713
Es/Ea | 1.9
Hs/Ea [mA/V] | 2.3
Sc/Ea2 [mA/V] | 0.28

## Slide 6: Gradient in 24 regular cells
Number of regular cells: Nc | 24
Bunch population: N | 3.72x109
Number of bunches: Nb | 312
Bunch separation: Ncycl | 6 rf cycles
Average unloaded of 100 MV/m
Average loaded of 100 MV/m

## Slide 7: Simulation setup 1 & 2
HFSS-quad
HFSS simulation of ¼ of the structure:
Surface conductivity (Cu): 58e6 S/m
Surface approximation: ds=6 μm,
Total number of tetrahedra: Ntetr = 1170348
Mesh density: ~ 40000 tetr / ¼ cell
S3P-quad
S3P simulation of ¼ of the structure made by Zenghai Li at SLAC-ACD
Surface conductivity (Cu): 57e6 S/m different
HFSS v10.1

## Slide 8: Simulation setup 3
HFSS-cells + couplers
HFSS simulation of ¼ of input coupler + segment of 5 deg. of the cells + ¼ of output coupler :
Surface conductivity (Cu): 58e6 S/m
Surface approximation: ds=1 μm,
Total number of tetrahedra for cells: Ntetr = 506075
Mesh density: ~ 350000 tetr / ¼ cell
(10 times higher than in HFSS-quad setup)
HFSS v11.1

## Slide 9: Measurement setup 1 (reflections)
Perfect
load
Perfect
load
VNA
port1
VNA
port2
1
2
3
4
Reflection = S11+S12
Frequency correction due to air (ε = 1.00059)
Frequency in vacuum: f’ = f*sqrt(1.00059) ≈ f + 3.5 MHz @ X-band

## Slide 10: Reflection: comparison
There is very small (~1MHz) or no difference in frequency between simulations and the air corrected measurements

## Slide 11: Measurement setup 2 (transmission)
1
2
3
4
Perfect load
Perfect load
VNA port1
VNA port2
Transmission = S13+S23
1
2
3
4
Perfect load
VNA port2
VNA port1
Perfect load

## Slide 12: Transmission: comparison
There is small difference between HFSS and S3P simulation results due to different conductivity used in the simulations
Both simulation results show higher transmission than air corrected measurements by about 0.2 dB

## Slide 13: Some useful equations
On the other hand
Finally

## Slide 14: Transmission: comparing group delay
There is no difference between both HFSS and S3P simulations and the air corrected measurements

## Slide 15: Transmission: comparing Q-factor
There is no difference in Q-factor (~7000) between HFSS and S3P simulations.
The measured Q-factor of about 6600 is lower than the simulated value by about 6 %.

## Slide 16: Measurement setup 3 (bead pull)
Perfect
load
Perfect
load
VNA
port1
VNA
port2
1
2
3
4
E2 ~ S11-<S11>

## Slide 17: Bead pull in complex plane
Measurements: S11-<S11>
HFSS-cells: E

## Slide 18: Bead pull: field magnitude
Measurements: |sqrt(S11-<S11>)|
HFSS-cells: |E|

## Slide 19: Field distribution at different frequencies
120o rf phase advance per cell frequency
11.424 GHz
Best match frequency
11.420 GHz

## Slide 20: RF phase advance per cell at different frequencies
120o rf phase advance per cell frequency: 11.424 GHz
Best match frequency:
11.420 GHz

## Slide 21: Summary table for T24_vg1.8_disk
 | S12 | Pout/Pin | τ=0.5ln (Pout/Pin) | tf=dφ/dω [ns] | tf=W0/Pin [ns] | Q=ωtf/2τ | Q=ωW /(Pin-Pout) | Q=ωW /Ploss | Pin100 24 cells [MW]
Measurements (Vac) | 0.715 | 0.511 | 0.336 | 60.9 | - | 6600 | - | - | 
HFSS ¼ | 0.732 | 0.536 | 0.312 | - | 58.4 | 6720 | 7010 | 7010 | 
HFSS InCoup OutCoup Cells 24+2 | 0.731 0.99875 0.99889 0.733 | 0.534 0.9975 0.9978 0.537 | 0.313 0.0013 0.0011 0.311 | 60.7 <- <- 59.9 | 59.8 0.46 0.36 59.0 | 6960 12700 11800 6910 | 6940 13000 12300 6880 | 6640 13200 12300 6580 | 42.2
3-cells model |  | 0.569- 0.533 | 0.282- 0.315 | - | 54.8- 59.37 | 6974- 6764 | 6980 |  | 41.1
S3P (SLAC) | 0.730 | 0.533 | 0.315 | 60.8 |  | 6927(σ=57e6) 6988(σ=58e6) |  |  | 42.4

## Slide 22: Summary on comparison
Simulation results of HFSS and S3P show very good agreement
Q-factor
  All 3 different ways of calculating Q-factor: S3P, HFSS-S-parameter solver and HFSS-eigenmode solver give very close values of about 7000
  The measurements of the T24_vg1.8_disk structure made at CERN show 6 % lower Q-factor of 6600
RF phase advance
  Both HFSS and S3P simulations and air corrected measurements simulations show 120o rf phase advance frequency of 11.424 GHz which is the design frequency
  Good agreement (± 0.5 MHz) between simulations and measurements demonstrates extremely high (sub-micron) precision of machining. For example, it is equivalent to ± 0.6 μm tolerance on the outer wall radius RF phase advance
Structure matching
  There is a design error in structure matching of about 4 MHz. The match frequency is 11.420 GHz . To some extent this is also in agreement with the measurements

## Slide 23: HFSS S-par solver: v10.1 versus v11.1

## Slide 24: Simulation setup 4
HFSS-cells + couplers
HFSS simulation of ¼ of input coupler
+ segment of the cells made in HFSS by 5 deg. Sweep (3D geometry created in HFSS)
+ ¼ of output coupler :
Surface conductivity (Cu): 58e6 S/m
Surface approximation: ds=1 μm,
HFSS v11.1
HFSS v10.1
versus

## Slide 25: Reflection
4 MHz
4 MHz

## Slide 26: Transmission
4 MHz

## Slide 27: Group delay and Q-factor

## Slide 28: RF phase advance per cell
HFSS v10.1
HFSS v11.1

## Slide 29: Summary on HFSS
It seems that there is a bug in HFSS v11.
The results of the simulations using HFSS v11 are shifted up in frequency by 4 MHz with respect to the results of the simulations using HFSS v10 or S3P
Use HFSS-S-parameter solver VERSION 10 or S3P